Performance of the Hydroclave for Decontamination of Biomedical Waste

Trials conducted on unit installed at Kingston General Hospital

Submitted to

Hydroclave Systems Corp.

1361 Middle Road, Kingston, ON.

&  Ontario Ministry of Health

Susan Springthorpe & Syed Sattar

Department of Microbiology & Immunology

Faculty of Medicine, University of Ottawa

Ottawa, ON.

November 1995

2

INDRODUCTION

Mr. Richard Vanderwal of Hydroclave Systems Inc. approached the University of Ottawa regarding testing of the machine (Hydroclave) which Hydroclave Systems Inc. had developed for the treatment of biomedical waste.  A protocol for such testing was submitted to the Ontario Ministry of Health on September 27, 1995.  Subsequent to verbal approval of this protocol given to Mr. Vanderwal, and authorization for him to conduct the testing using normal hospital waste, testing was conducted by the University of Ottawa on a prototype machine installed at Kingston General Hospital between October 4, 1995 and November 11, 1995.  This report describes these tests in detail.

METHODOLOGY

The Hydroclave installed at Kingston General Hospital is shown in Figure 1.  As noted above, the protocol for testing this machine was submitted to the Ontario Ministry of Health for approval in September 1995.  Tests were conducted according to the submitted protocol.  The only change was due to an unfortunate incident, which occurred when staff at Kingston General Hospital accidentally discarded all the biological indicator which were stored in the cold room.  The tests were therefore conducted in two stages.  The original tests to determine cycle time (Runs 1-6) and one additional run at 121°C (Run 8) and one run at 132°C (Run 7) were conducted using all of the biological indicators originally described.  Discard of the biological indicators occurred shortly after October 6, 1995.  This meant that an additional batch of biological indicators with nominal population in excess of 106 .  Sportorol STS-06 were obtained with a nominal population of 1.7 X 106 (Appendix 1) and these were used for the remaining tests at 121ºC and 132ºC which took place between October 23 and November 11, 1995.  All biological indicators were held in a specially constructed container (Figure 2) in order to hold them away from the rotating blades of the Hydroclave and to prevent them from damage during the operation of the unit.

It was difficult to assess the water content of the waste loads from the hospital but additional water (20 LB) was added to some of the loads as indicated in the tables in the results section to make sure the load was wet.  Because of the way the Hydroclave functions however, the water content of the loads is used to generate the pressure and so the water content is a less important factor than in a conventional autoclave.  Whenever steam was used in addition to the water content of the load, this is shown in the tables of results.

As described in the original protocol, measurements of temperature and pressure during the runs were taken manually during this initial testing of the prototype.  The timing for the run was started only after the proper temperature was reached.  Subsequently, automated instrumentation will be installed for this purpose.  At that time, the cycle time will start only after the thermocouple indicates the proper preset temperature.  The thermocouple will be located at the coolest part of the unit, which will be where any water in the load accumulates at the base of the machine.

Waste load weight reductions were determined on two runs only.

3

RESULTS & DISCUSSION

Specifications for the Bacillus stearothermophilus spores in the biological indicators used are shown in Table 1; all BI's were stored according to the manufacturer's recommendations.  Populations were determined for all the biological indicators according to USP procedures.  These values are shown in the last column: two of the BI's were marginally out of compliance with USP guidelines.  Kilit showed values marginally below 50% of the nominal value, but this was not felt to be significant since it appeared to be the most robust during the testing.  The Sportrol STS-06 was marginally more than 300% of the nominal value.  Likewise, this was not felt to be significant.  More disturbing was the lack of any growth in the Attest even after the minimum cycle time (Table 2) in spite of the fact that determination of the mean titre showed that it was in compliance with its nominal value.  We have no explanation for this and could not examine it further since it was among the biological indicators discarded.

Table 2 shows a summary of all 16 runs performed.  Runs 1-6 determined the cycle time at 121ºC to be 30 min. Run 7 confirmed that an equivalent result could be obtained at 132ºC in only 15 min. Run 8 repeated the 30 min. run at 121ºC.  After the loss of the biological indicators and their replacement by Sportrol STS-06 with a nominal population of 1.7 X 106, runs 9-11 repeated the 30 minute cycle at 121ºC and confirmed it to be effective.  Runs 12-14 repeated the 15-minute cycle at 132ºC, which equally was shown to be effective at inactivating the biological indicators.  Runs 15 and 16 were selected because they were comprised entirely of garburator waste and were very heavy wet loads.  The biological indicators included in these runs at 132ºC for 15 min. at temperature showed no growth.  Tables 3-18 show details of the temperature and pressure achieved during the runs.  It can clearly be seen from these data that the different nature of the waste in each of the runs had an effect on the length of the entire run time, although the time for which the load was held at temperature was constant.  These differences in run time can be predicted from the way the Hydroclave is designed to function.

Weight reductions in the waste load during operation of the Hydroclave were determined on only runs 1 and 2; these showed reductions of 32.6% and 38.4%, respectively.  However, the weight reductions achievable will be a function of the percentage of water in the load.  Figure 3 and 4 show the dry waste removed from the Hydroclave after treatment.  The overall size of the waste can be seen from the cans and jugs it contains.  Addition of a grinder to further reduce the size of the waste and make it unrecognizable may be desirable when biomedical waste itself is being treated.

CONCLUDING REMARKS

The Hydroclave appears to be highly suitable for the treatment of biomedical waste and can achieve an inactivation of microbial load of  >106 equivalent of B. stearothermophilus within 30 minutes at 121ºC or within 15 minutes at 132ºC.  Although the cycle time may be able to be shortened even further at 132ºC, 15 minutes is a relatively practical period and a shorter exposure may be unnecessary.  If however it was desired to shorten the cycle time at temperature even further at 132ºC, then further timed tests would be required.

Note: Although the Kilit biological indicator appeared to be the most robust in the initial tests, no equivalent bioindicator with > 106 spores was available.

4

Table 1	Specifications of Bacillus stearothermophilus in Biological Indicators used

| Biological Indicator | Manufacturer | Lot # | ATCC Strain # | D121 [D132] (min) | Expiry date | Nominal population per strip/ampoule | Mean Titre per strip/ampoule |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Attest 1262/1262P | 3M | 282 | 7953 | 1.8[n/a] | July 97 | 4.60E+05 | 7.10E+05 |
| Kilit | BBL | L5GMOD | 7953 | 1.8[n/a] | June 96 | 4.40E+04 | 2.04E+04 |
| Sportrol STS-05 | NAMSA Inc. | S49503 | 7953 | 1.9[0.12] | Dec.96 | 1.90E+05 | 1.30E+05 |
| Biosign | MDT | 181195B | 7953 | 2.78[0.44] | Aug. 96 | 1.00E+04 | 2.86E+04 |
| In house BI | University of Ottawa | n/a | 12980 | n/a | not available, freshly prepared | not available | 6.00E+05 |
| In house BI | University of Ottawa | n/a | 12980 | n/a | not available, freshly prepared | not available | 6.00E+04 |
| In house BI | University of Ottawa | n/a | 12980 | n/a | not available, freshly prepared | not available | 6.00E+03 |
| In house BI | University of Ottawa | n/a | 12980 | n/a | not available, freshly prepared | not available | 6.00E+02 |
| Sportrol STS-06 | NAMSA Inc. | S48403 | 7953 | 2.1[0.1] | Feb. 97 | 1.7 X 106 | 6.00E+06 |

n/a = not available

5

Table 2

Determination of cycle time required for kill of Bacillus stearothermophilus and test runs at 121ºC and 132°C

| Run # | Time (min) | Temp (ºC) | Pressure (psi) | Attest (7.1 x 105) | Kilit (2 x 104) | Sportrol (1.3 x 105) | Biosign (2.9 x 104) | BI (6 x 105) | BI (6 x 104) | BI (6 x 103) | BI (6 x 102) | Sportrol (6 x 106) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | 15 | 121 | 20 | - | + | + | + | + | + | + | + | nd |
| 2 | 10 | 121 | 15 | - | + | + | + | + | + | + | + | nd |
| 3 | 20 | 121 | 22 | - | + | + | + | + | + | + | - | nd |
| 4 | 25 | 121 | 22 | - | + | - | - | - | - | - | - | nd |
| 5 | 30 | 121 | 25 | - | - | - | - | - | - | - | - | nd |
| 6 | 40 | 121 | 35 | - | - | - | - | - | - | - | - | nd |
| 7 | 15 | 132 | 45 | - | - | - | - | - | - | - | - | nd |
| 8 | 30 | 121 | 30 | - | - | - | - | - | - | - | - | nd |
| 9 | 30 | 121 | 30 | nd | nd | nd | nd | - | - | - | - | - |
| 10 | 30 | 121 | 28 | nd | nd | nd | nd | - | - | - | - | - |
| 11 | 30 | 121 | 30 | nd | nd | nd | nd | - | - | - | - | - |
| 12 | 15 | 132 | 40 | nd | nd | nd | nd | - | - | - | - | - |
| 13 | 15 | 132 | 37 | nd | nd | nd | nd | - | - | - | - | - |
| 14 | 15 | 132 | 42 | nd | nd | nd | nd | - | - | - | - | - |
| 15 | 15 | 132 | 37 | nd | nd | nd | nd | - | - | - | - | - |
| 16 | 15 | 132 | 37 | nd | nd | nd | nd | - | - | - | - | - |

nd = not done